## Slide 1
David Rach, Ph.D.
UMGCCC Flow Cytometry Shared Resource
Baltimore, MD
November 05, 2026
Slides
“Deciphering the black box”: leveraging shared-resource laboratory
resources to tackle spectral flow cytometry’s quirks

## Slide 2: Spectral is amazing

## Slide 3: Making analyses possible

## Slide 4: Success due to us as a community

## Slide 5: Still has its quirks

## Slide 6: Co-expression in unmixing

## Slide 7: Unmixing controls autofluorescence

## Slide 8: Unmixing controls single colors

## Slide 9: QC goes bad with bleach

## Slide 10: Mix art, science and hoodoo

## Slide 11: SRLs as force for good

## Slide 12: Wide range of panels and cell types

## Slide 13: Instruments over lifespan

## Slide 14: Who spots the issues

## Slide 15: Im tired maaannn

## Slide 16: Automate the detection

## Slide 17: And decipher some underlying unanswered

## Slide 18: Manufacturer accessability

## Slide 19: What we do as far as QC

## Slide 20: A little programming knowledge makes you dangerous

## Slide 21: Things you may have not known about your Aurora

## Slide 22: Sit Flushes in App Log

## Slide 23: Usage Plots vs. hour

## Slide 24: Commando Users

## Slide 25: Usage Pattern Over Time

## Slide 26: QC CSV folder

## Slide 27: Temperature

## Slide 28: QC Time

## Slide 29: Parse for errors, sequences

## Slide 30: FCS files undefined

## Slide 31: QC doing what is says undefined

## Slide 32: CSVs to dashboard undefined

## Slide 33: Damaged QC beads

## Slide 34: Expt files undefined

## Slide 35: Users not knowing about threshold

## Slide 36: Correct antibody

## Slide 37: Stain index and prediction

## Slide 38: Correlating variation

## Slide 39: Is my unstained zombied?

## Slide 40: Future works in progress

## Slide 41: Regardless, what one person can do

## Slide 42: But what if we had a community soapbox

## Slide 43: Manufacturers can help or get out of the way

## Slide 44: Rise together stronger

## Slide 45: Open-source and shared

## Slide 46: R not just for the analyst

## Slide 47: Cytometry in R slides

## Slide 48: My scripts are here

## Slide 49: Call for Cytometry QC data

## Slide 50: Acknowledgements

## Slide 51: Questions

## Slide 52: Spectral Flow Cytometry is amazing!!!
Capable of resolving remarkably similar fluorophores → increased number of markers per panel compared to conventional flow cytometry (CFC)
Allows for rapid acquisition of millions of events in a few minutes → typically greater number of acquired cells per sample compared to mass cytometry (MC)
Spectral Flow Cytometry (SFC) datasets have an unique combination of width and depth compared to preceding technologies

## Slide 53: Spectral Flow Cytometry is amazing!!!
Expand our knowledge about immune system heterogeneity
Able to efficiently acquire greater number of cells → carry out a more thorough profiling of rare cell subsets
Build one panel, evaluate multiple cell types → invaluable for clinical studies with limited biospecimen

## Slide 54
Adapted algorithmic tools struggle with unmixing errors
Adapted software packages heavily rely on having a consistent median fluorescent intensity (MFI) to cluster cells based on positive or negative marker expression
Unmixing errors can seriously shift the MFI ranges for “negative” and “positive” cells for individual acquired specimens
Existing batch correction / normalization / scaling tools target instrumental or experimental level variation for correction

## Slide 55
“We discovered a novel dendritic cell cluster”

## Slide 56
Building open-source tool-sets to address SFC’s quirks
Spectral Signature Characterization and Isolation
Automated Instrument
Quality-Control Dashboards
Incorporation of manually-annotated gating information
Rust

## Slide 57: Automated Daily Processing
Behind-the-scenes:
Using R (and increasingly Rust) to process the .fcs files
Our Quarto Website gets updated twice daily via a GitHub action
Automated transfer to GitHub via Windows Task Scheduler

## Slide 58: Instrument Quality Control Dashboard
Dashboard
How-To

## Slide 59: Instrument Quality Control Dashboard

## Slide 60
PE-Vio 770
APC-Vio 770
Spectral Unmixing

## Slide 61
PE-Vio 770
APC-Vio 770
Mismatched references lead to errors in attribution

## Slide 62
How to identify problematic unmixing controls
Despite the critical importance unmixing controls play, few tools permit us to determine quality of the individual control before unmixing
Existing quality control algorithms (PeacoQC, FlowAI, etc.) primarily focus on detecting laser and fluidic issues in already acquired .fcs files.
Unmixing errors can be introduced by tandem degradational, additional autofluorescences, undiagnosed instrumental issues, etc.

## Slide 63: Luciernaga
Code

## Slide 64: Luciernaga
Code

## Slide 65: Luciernaga
Code

## Slide 66: Automated tools for screening unmixing controls
Autofluorescence Isolation
Residual Autofluorescence
APC-Fire 810
APC
Tandem Degradation

## Slide 67
When the fluorophore degrades, noticeable increase in variant signatures, making up substantial proportion of all cells in the unmixing control
Automated screening all acquired unmixing controls from campus users during off-hours
Identifying degrading fluorophores
Fluorophores in good shape → one predominant normalized signature within the unmixing control raw .fcs file

## Slide 68
This approach also works to isolate variant autofluorescence signatures from unstained controls
Return directly as an .fcs file containing just cells with same signature for use as an autofluorescent tag in SpectroFlo
Unlike other approaches, the isolated cells share normalized signatures, only differing in terms of brightness
Isolating additional autofluorescence signatures

## Slide 69
“As bright or brighter”
Identifying and switching these dim cell single-color controls to bead single-color controls significantly improved unmixing for our panels
Dim antigen-fluorophore combinations acquired as cell single-color controls often have residual autofluorescence contribution to their signature
Residual autofluorescence in dim cell single-color controls
Signature with
residual AF
Signature without
residual AF

## Slide 70: What good is any tool if the community is not able to use it?
Few resources to learn Cytometry in R to begin with
Those that exist aimed at those with intermediate-advanced R skills, not beginners
Know-how concentrated at a few institutions
Most self-learning attempts end in frustration

## Slide 71: Cytometry in R: A free weekly course for flow cytometrist with no-to-little coding experience
Website

## Slide 72: YouTube
Three live-streams a week (accommodating different timezones and schedules)
Recordings available immediately after
Best of the weekly livestreams added to the playlist, cutting down on editing

## Slide 73: Website
Created with Quarto, hosted via GitHub pages.
Code, code outputs, screenshots and explanations
Searchable
New weekly material added as a new .qmd file

## Slide 74: GitHub
Code and datasets available via the GitHub repository
Course participants start course by creating GitHub account and forking the repository
After Positron is linked to GitHub account, pull locally
When new material posted each week, sync and pull

## Slide 75: Discussion Forum
GitHub Discussions page has been repurposed as a community forum for questions, etc.

## Slide 76: Half-way through this passion project
Wrapping up Cytometry Core module
Gathering steam to start on the Unsupervised Analysis section

## Slide 77: As scientist, our community is our strength

## Slide 78: Just my two over-caffeinated cents

## Slide 79: Acknowledgments
Dr. Cristiana Cairo
Dr. Kirsten E. Lyke
Dr. Xiaoxuan Fan
All mothers and their infants, and their communities
All the open-source contributors, developers, and maintainers, that collectively made the tool sets without which all of this would not be feasible
Cytometry
Discord
Bioconductor
Zulip
Dr. Daniela Franco
Dr. Natarajan Ayithan
U01 HD092309
This publication was supported by funds through Maryland Department of Health’s Cigarette Restitution Fund Program – CH- 649-CRF and the National Cancer Institute – Cancer Center Support Grant (CCSG) – P30CA134274.d Service
www.linkedin.com/in/david-rach/
Want to collaborate? Reach out!
Hiring?

## Slide 80: Questions?
UMGCCC FCSR
InstrumentQC dashboard
YouTube @CytometryInR
Website - CytometryInR
Luciernaga
Coereba
InstrumentQC How-To